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Overview：
RQ810G is composed of RC810G dissolution tester connected to

QY8 sampler. It has electric synchronous medicine, single sampling amount of 100ml and single collection amount of 15ml. It can perform sampling and filtration with high precision. Its performance indicators are in full compliance with the requirements of the Chinese Pharmacopoeia (2015 edition) and the "Guidelines for Mechanical Calibration of Dissolution Testers".
Function parameter：
It can realize the test of basket method, paddle method, small cup method 1, dish method (large and small dish) 1, rotating tube method (long and short tube) 1 to meet blank and control requirements.。

 Machine head electric lift

Automatic constant clutch

Simultaneous administration and sequence administration

Two-color lighting system
Use a precision glass syringe as the sampling power.
Adopt high precision, high pressure and high chemical stability valve body.
Built-in micro printer.
Multi-user management。
10.WiFi, Bluetooth, RS232, RS485 expansion interface。

11.Biometric identification。
12.Highly chemically stable polytetrafluoroethylene tubing。
13.Filter discharge function to reduce adsorption.

14.With equal amount of rehydration。
15.With independent pipeline system, it can meet the control 16.requirements of blank and standard for dissolution test.A cylindrical filter head having a pore size of 10 μm was sampled, and a φ25 mm syringe filter having a pore size of 0.8 μm/0.45 μm/0.22 μm was selected as needed for the secondary filtration.
17.A 16-seat test tube rack can hold up to 128 test tubes.
18.The sampling tube is equipped with a silicone cap to prevent evaporation and contamination of the sample.

19.Run online with the dissolution apparatus。
20.Ampoule collection device。
21.High sensitivity capacitive touch screen。
22.authority management。
23.Audit trail。
Technical indicators:

RC810G Dissolution tester

	﻿Model
	RC810G

	Number of experimental devices
	8

	Water bath temperature
	Room temperature—45℃

	Temperature resolution
	0.01℃

	Stirring speed
	25RPM—250RPM

	Stirring speed resolution
	0.1RPM

	Dosing channel
	8

	Head lift
	electric

	Lighting system
	Red light, white light

	amount of users
	7users

	user level
	Third level

	User program
	99

	Record saving
	128 *366 days


QY8 Sampler
	﻿Model
	QY8

	Sampling channel
	8

	Number of samples
	16

	Longest sampling interval
	99 hours 59 minutes

	Minimum sampling interval
	1 minute for the first time, the remaining 2-10 minutes (depending on the sample size, sampling method and whether it is rehydration)

	Continuous working time
	99 hours 59 minutes × 16

	Sampling filter cycle
	<30 seconds

	Sample volume (or fluid volume)
	1ml—100ml

	Single tube distribution
	0ml—15ml

	Sampling accuracy
	RSD1%（10ml）

	User program
	99

	amount of users
	7 users

	user level
	Third level

	Record saving
	128 *366 days
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